| Item No | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | CARBON/SULFUR ANALYZER | Features Long MTBM (Mean Time Between Maintenance) High Throughput Fast & Easy Maintenance Advanced Operation software Application Quantitative analysis of carbon contained in bone substitute material Quantitative analysis of carbon and sulfur in tires Measurement of Carbon and Sulfur in Cement Battery Evaluation in Electrical Equipment Technical specifications Required sample amount: 1g ± 0.1g Typical Meas. Time: 70 seconds/cycle Carbon: Meas. Range (m/m)-1.6 ppm - 6.0% Blank accuracy σn-1 ≦ 0.8 ppm Sulfur: Meas. Range (m/m) 2.0 ppm - 1.0% Blank accuracy σn-1 ≦ 1.0 ppm Sample meas. Accuracy: Carbon σn-1 ≦ 2.0 ppm or RSD≦1.0% Sulfur σn-1 ≦ 2.0 ppm or RSD≦1.5% Utility: Carrier gas Oxygen (Purity; 99.5%, Pressure; 0.30 0.33MPa) Operation gas Nitrogen (Purity; 99.5%, Pressure; 0.35 0.38MPa) Power: 200/220/240V, 50/60Hz, 5kVA Dimensions Main unit* 500(W) x 725(D) x 710(H) mm Mass Main unit approximately 134 kg Data processing and operation: USB data communication with PC Windows 8.1 Touch panel or Key board and mouse operation Included Articles Accessories - Crucible Pre-heating Unit - Autosampler Consumables - Ceramic crucible - Accelerators Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |