| Item No | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | CARBON/SULFUR ANALYZER | Features Long MTBM (Mean Time Between Maintenance) High Throughput Fast & Easy Maintenance Advanced Operation software Application Quantitative analysis of carbon contained in bone substitute material Measurement of Carbon and Sulfur in Cement Battery Evaluation in Electrical Equipment Measurement of Sulfur in Rubber Technical specification Required sample amount: 1g ± 0.1g Typical Meas. Time: 70 seconds/cycle Carbon Meas. Range (m/m) 0.6 ppm - 10.0% Blank accuracy σn-1≦ 0.3 ppm Sulfur Meas. Range (m/m) 0.6 ppm - 1.0% Blank accuracy σn-1 ≦ 0.3 ppm Sample meas. Accuracy Carbon σn-1 ≦ 0.3 ppm(less than 20 ppm) σn-1 ≦ 1.0 ppm or RSD≦0.5%(more than 20 ppm) Sulfur σn-1 ≦ 0.3 ppm(less than 20 ppm) σn-1 ≦ 1.0 ppm or RSD≦0.75%(more than 20 ppm) Utility Carrier gas: Oxygen (Purity; 99.5%, Pressure; 0.3-0.33 MPa) Operation gas: Nitrogen(Purity; 99.5%, Pressure; 0.35-0.38 MPa) Power: 200/220/240V, 50/60Hz, 5kVA Dimensions Main unit*500(W) x 725(D) x 710(H) mm Purifier unit 160(W) x 560(D) x 512(H) mm Mass Main unit approximately 134 kg Purifier unit approximately 21 kg Data processing and operation: USB data communication with PC Windows 8.1 Touch panel or Key board and mouse operation Included Articles Accessories - Crucible Pre-heating Unit - Autosampler Consumables - Ceramic crucible - Accelerators Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |