| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | THERMOGRAVIMETRIC ANALYZER | Features - Measurement of up to 19 samples in one analysis - Sample weights of up to 5 g - Fast heating rates, accurate temperature control - High-performance, precise weighing cell Application Analysis of coal & coke Analysis of building materials Analysis of food Technical specification Measured elements: ash content, moisture, volatiles Samples: inorganic, organic, synthetic Field of application: agriculture, biology, chemistry / plastics, coal / power plant, construction materials, environment / recycling, food, geology / mining, glass / ceramics, medicine / pharmaceuticals Furnace: resistance heated ceramic furnace, programmable in 1 °C steps from 50 °C up to 1000 °C Detection method: balance Max. number of samples: 19 crucibles + 1 reference crucible Balance resolution: 0.0001 g Balance precision: 0.02 % RSD oxygen 99.9 % pure (2 - 4 bar) and,or nitrogen 99.9 % pure (2 - 4 bar) and air 99.5 % (5 - 6 bar) Power requirements: 230 V, 50/60 Hz, max. 32 A Dimensions (W x H x D): 55 x 52 x 62 cm Weight~ 65 kg Required equipment: PC, monitor, external exhaust (exhaust diameter: 100 mm / blower with 4m-3/min) Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |