| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | AMBIENT CARBON MONOXIDE MONITOR | Description A device for continuously monitoring CO concentrations using a non-dispersion cross modulation infrared analysis method Features Cross-flow non-dispersive infrared detector; Low range: 0-5 ppm F.S. • AS type (anti-shock) interference-compensating detector and purifier for reference gas. • Reference gas purifier oxidizes CO to CO2 to eliminate interferences. • Optics remain free of foreign matter with elimination of reflecting mirrors Technical specification Principle: Cross flow modulation, non-dispersive infrared (NDIR) absorption technology Application: CO in ambient air Range: Standard ranges: 0-10/20/50/100 ppm; 0-5/10/20/50 ppm; auto range ~ manual range selectable; can be operated by remote switching. Optional (measurable) ranges: 4 ranges selectable from 0-200 ppm, within 10 times range ratio; auto range ~ manual range selectable; can be operated by remote switching. Repeatability: ±1.0% of F.S. Linearity: ±1.0% of F.S. Sample gas flow rate: Approx. 1.5L/min Indication: Measured value, range, alarm, maintenance screen Alarms: During AIC, zero calibration error, span calibration error, temperature error in catalyzer, etc. On-screen messages are available in four languages: English, German, French, and Japanese Input/output • 0-1 V/0-10 V/4-20 mA, to be specified (2 systems: either (1) momentary value and integrated or (2) moving average value) • Contact input/output • RS-232C Power: 100/110/115/120/220/230/240 VAC, 50/60 Hz (to be specified) Dimensions: 430(W) x 550(D) x 221(H) mm Mass: Approx. 16 kg Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |