| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | AMBIENT SULFUR DIOXIDE MONITOR | Description A device for the continuous monitoring of atmospheric SO2 using UV fluorescence Features • UV-fluorescent detector resistant to moisture interference; Low range: 0-0.05 ppm F.S. • Alternately to FDP measurements, the detector design is highly SO2 selective and requires no supplemental gas. • Integral HC-cutter with a selective membrane eliminates interfering components. • Incorporates a lamp intensity compensator. • Uses a built-in inert PTFE sample inlet filter. Technical specification Principle: UV fluorescence (UVF) Application: SO2 in ambient air Range: Standard ranges: 0-0.05/0.1/0.2/0.5 ppm; auto range ~ manual range selectable; can be operated by remote switching. Optional (measurable) ranges: 4 ranges selectable from 0-10 ppm, within 10 times range ratio; auto range ~ manual range selectable; can be operated by remote switching. Repeatability: ±1.0% of F.S. Linearity: ±1.0% of F.S. Sample gas flow rate: Approx. 0.7L/min Indication: Measured value, range, alarm, maintenance screen Alarms: During AIC, zero calibration error, span calibration error, temperature error in catalyzer, etc. On-screen messages are available in four languages: English, German, French, and Japanese. Input/output • 0-1 V/0-10 V/4-20 mA, to be specified (2 systems: either (1) momentary value and integrated or (2) moving average value) • Contact input/output • RS-232C Power: 100/110/115/120/220/230/240 VAC, 50/60 Hz (to be specified) Dimensions: 430(W) x 550(D) x 221(H) mm Mass: Approx. 19 kg. Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |