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	AMBIENT HYDROGEN SULFIDE MONITOR
	Description
Ambient H2S monitor measures SO2 converted H2S through oxidation catalyst based on Ultra Violet Fluorescence. In order to reduce measurement error by water concentration fluctuation in ambient, humidifier is set at sample line. It enables long-term stability according to stabilizing catalytic reaction.
Features
Monitor H2S with high sensitivity range in ambient air continuously.
Set humidifier at sample line to stabilize catalytic reaction.
Technical specification
Measurement component: Hydrogen sulfide (H2S) in ambient air
Measurement principle: Oxidation catalyst + Ultra Violet Fluorescence
Measurement range: 0-0.1/0.2/0.5/1.0 ppm
Lower detectable limit: 2 ppb
Repeatability: ±3.0% of F.S.
Linearity: ±2.0% of F.S.
Zero drift: ±2% of F.S./day
Span drift: ±5% of F.S./day
Response time (T90): 180 sec. or less (from CU-1converter unit)
Sample gas flow rate: approx. 0.7L/min
Indication: Measuring value, alarm
Inpt/Output: 0-1V/0-10V/4-20mA (2 systems: either (1) momentary value and integrated of (2) moving average value)
Contact input/output
RS-232C (option)
Temperature/Humidity	0-40°C / 85% or less
Power: 100/110/115/120/220/230/240 VAC, 50/60Hz
Documentation:
· CE marking and declaration conformity
· Documents for the Installation and Operational Qualification (IQ / OQ)
· Functional Specification - Design Qualification (DQ) and others relevant document
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