| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | TRACE GAS MONITOR | Description For monitoring quality control of manufactured high-purity gases at air separation and semiconductor plants Application Test Gas Generation in Research and Testing Laboratories Low Concentration Gas Monitoring in Coal and Consumable Fuels Features Ultra-high Sensitive Measurements - A cross-flow modulation dual-beam non-dispersive infrared (NDIR) analyzer provides continuous and stable zero drift-free measurements. - Minimum detection limit (MDL) of 10ppb which can support applications where high accuracy measurements are required. Measurement of trace impurities in a balance gas - Representative gases such as N2, O2, He, Ar, H2 and Air are supported. Easy maintenance-free operation - Operator friendly screen menus simplify analyzer operation, calibration and measurements - No optical alignment ever required. - A touch color LCD display allows operators to view graphs of accumulated data. Technical specification Gases measured: CO, CO2, CH4 Bulk gases: N2, O2, He, Ar, H2, Air Number of measurement components: 1 component or 2 components (specify a measuring balance gas) Measuring principle: Cross-flow modulation non-dispersive infrared (NDIR) absorption method Measurement ranges: 0 to 1/2/5/10 ppm Lowest detection limit (2σ): 10 ppb Repeatability: ≤ ± 2% of full scale Indication error : ≤ ± 2% of full scale Zero Drift: ≤ ± 0.02 ppm/day, ≤ ± 0.03 ppm/week Span Drift: ≤ ± 2% of full scale/day, ≤ ± 2% of full scale/week Response time (T90): ≤ 180s Gas flow rate: Sample gas : approx. 3.5L/min, reference gas : approx. 3.5L/min, span gas: approx. 3.5L/min Note: A delivery pressure of 50-100kPa is recommended for both sample and reference gas. Analog output: Max. 2 channels isolated output (2 components) Select one from 0 to 1V, 0 to 10V, 0 to 16mA, 4 to 20mA or 0 to 20mA. Current output : Allowable load resistance 750Ω or less Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |