| Item No | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | HYDROGEN ANALYZER | Description Designed for the rapid and accurate determination of hydrogen in these materials. Which uses the heat extraction technique and is equipped with a resistance furnace with quartz tube. The temperature can be set up to 1000 °C in steps of 1 °C. APPLICATION EXAMPLES alloys, copper, steel Features and Technical specifications - hydrogen determination with heat extraction technique for determination of residual hydrogen - high-capacity thermal conductivity cell - easy calibration with standards or gas dosing - precise measurements even for low concentrations - for samples of up to 10 g and 0.8 x 6 cm size - powerful software (multilingual, customized display, export of results) - low maintenance - robust design allows usage in production control and laboratory Measured elements: hydrogen Samples: metals, steel Furnace alignment: horizontal / tilting function Field of application: engineering / electronics, glass / ceramics, steel / metallurgy Furnace: resistance furnace with quartz tube, adjustable up to 1000 °C (operating temperature 900 - 1000 °C) Detection method: thermal conductivity Maximum sample size: ø 13 x 60 mm Typical analysis time: 3 - 15 min Chemicals required: Schuetze reagents, magnesium perchlorate, sodium hydroxide Gas required: nitrogen 99.995 % pure (2 - 4 bar / 30 - 60 psi) Power requirements: 230 V, 50/60 Hz, 2 A, max. 450 W Dimensions (W x H x D) :75 x 52 x 60 cm Weight~ 40 kg Required equipment: PC, monitor, balance (resolution 0.0001g) Included Articles Optional accessories: voltage stabilizer 5 KVA Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |