| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | DATALOGGER IO-EXPANDER | Features - Acquisition of raw values and control of analyzers (analogue, serial, Ethernet) - Automatic and manual controlled calibration of analyzers - Generation of two different average values - Save data storage of calibration-, average values and messages - Data communication with central location (modem, LAN) - Automatic sending of event message (failure of analyzer, exceeding of limited values…) as SMS, e-mail or to the central location - Integrated web interface enables access to values of station - Hardware without moving parts (passive cooling) - Configurations- and visualization software IOVIS lite saved on the integrated Samba server Technical specification Hardware: Embedded PC components PC/104 Operating system Linux Database: 8640 short term average (1 - 60min) 8640 long term average (5min - 24h) 1200 calibration data 400 messages Data request, configuration: Web Interface IOVIS Central software Interfaces Ethernet 10/100 MBit, TCP/IP 8+2 RS232 (expandable to 18) 16 analogue inputs, 0(4) - 1(10)V, 16bit 8 digital inputs (optocoupler expandable to 40/72) 4 relay outputs contacts (expandable to 20/36) Ambient temperature: 0 - 40°C Power: 110 - 230VAC 50/60Hz ca.30VA Dimensions: 449(W) x 315(D) x 133(H) mm Case: 19inch case, 84HP, 3U Mass: Approx. 5kg Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |