| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | AUTOMATIC COD MONITOR | Features - Uses only 1/10 the reagent of the existing systems unique dispensing and quantification device enable it to maintain measurement accuracy while using only 1/10 the quantity of reagent required. A reagent delivery service also eliminates the need for time-consuming reagent preparation. This reduces the time and trouble involved in reagent replacement and cuts running costs in half. - Touch panel LCD makes operation easier The use of a color LCD with a touch panel makes operation intuitive, and lower power consumption reduces environmental load. These and other functions attuned to the times will enable the system to meet future water quality management needs. - Cuts power consumption in half - Direct heating method requires no cooling water - Full range of interface options (RS-232C, RS-485, USB, and MODBUS®) - 2 Lines for Different Types of Samples and Applications Technical specification Object: COD concentration in water Dimensions : 600 (W) mm × 510 (D) mm × 1600 (H) mm*1 Mass Approx. 150 kg Power source: 100 V to 240 V AC ±10%, 50 Hz/60 Hz Power consumption 100 V to 120 V AC: Approx. 250 VA 120 V to 240 V AC: Approx. 350 VA Performance Measurement range: 0 mg/L to 20 mg/L (Actual upper limit of measurement is half-scale of the range.) 0 mg/L to 30/40/50/100/200/500/1000/2000 mg/L (1-dilution type) Reproducibility (with standard solution for glucose) 20 mg/L: Within ±1% of full-scale value 30 mg/L to 500 mg/L: Within ±2% of full-scale value 1000 mg/L to 2000 mg/L: Within ±5% of full-scale value Heating method: Direct heating End point detection: Potentiometric titration at constant current Titration method: Micro syringe titration Measurement interval: 60 minutes Measuring conditions Ambient temperature: 2℃ to 40℃ Ambient humidity: Relative humidity: 85% max. (without condensation) Power supply voltage fluctuations: 100 V to 240 V AC ±10% Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |