| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | ORGANIC POLLUTANT MONITOR | Features Sensor that makes full use of Rotary Cell Length Modulation - Zero drift correction is performed in every cycle of measurement. This correction can eliminate the effect of any interference on cells. - Cell length modulation provides readings from multiple measurements, making available results from various cell lengths. This allows the single device to make measurements on a large range of concentrations from as low as 0 to 0.1 Abs, up to 0 to 5.0 Abs. Easy-to-use Control Panel An interactive setup screen eliminates any doubts that a user may have while operating the unit. The Converted COD and TURB Display allow readings to be confirmed on site. Technical specification Measurement item: UV absorbance, VIS absorbance, COD conversion concentration, TURB (turbidity) conversion concentration Measurement principle Dual optical path, Dual wavelength, rotary cell length modulation method Measurement wavelength: Ultraviolet: 253.7 nm, Visible light: 546.1 nm Structure of analyzer unit: Circulation type Measurement range (converted by cell 10mm) UV absorbance/VIS absorbance Full scale: 0.1 Abs to 5.0 Abs (setting by 0.1 Abs is possible) Minimum scale 0.0001 Abs (changeable to 0.001 Abs) Repeatability: Within ±2.0% of full scale (Within ±5.0% of full scale for 2.6 Abs to 5.0 Abs) Linearity: Within ±2.0% of full scale (Within ±5.0% of full scale for 2.6 Abs to 5.0 Abs) Stability: Within ±2.0%/24 h of full scale (Within ±4.0%/24 h of full scale for 2.6 Abs to 5.0 Abs) Response time T90, within 1 minute (flow rate of approximately 5 L/min) Cleaning method (cleaning cycle): Automatic cleaning by wiper (automatic continuous cleaning) Display method LCD display: 320 × 240 monochrome graphics LCD data display with dot backlighting (a touch-panel) Display content: UV absorbance, VIS absorbance, UV-VIS absorbance, COD conversion concentration, TURB (turbidity) conversion concentration Ambient temperature and humidity: Ambient temperature: 0°C to 40°C, Ambient humidity: 85% or below Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |