| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | PANEL-MOUNT TYPE CONDUCTIVITY METER (FOUR-WIRE TRANSMISSION, FOR WIDE RANGE) | Features - Simultaneous 4-channel output Advanced concentration management is possible through simultaneous 3-channel output of low, middle, and high concentrations, etc. - Five contact alarms - Input signal can apply for analog output range selection - Including nitric acid, phosphoric acid, sodium hydroxide, and 2 user-specified types, the concentration of a total of 5 types of chemical solutions can be output - Built-in Communications Function Technical specification Measuring method: Four-electrode type Cell constant: 0.1/cm Temperature sensor specifications: Resistance thermometer bulb 1000 Ω/0℃ Measuring range: Conductivity: 0 mS/cm to 200 mS/cm (measuring range of conductivity before temperature compensation: 0 mS/cm to 500 mS/cm) Temperature: 0.0℃ to 100.0℃ Display and output range: Conductivity: 0 mS/cm to 2000 mS/cm Temperature: −30℃ to 160℃ Concentration: 0% to 5%: NaOH, HNO3, H3PO4 0% to 100%: Concentration of the component Temperature compensation range: 0℃ to 100℃ (a temperature less than 0℃ and more than 100℃ is calculated as an extension of the range) Ambient temperature: −5℃ to 55℃ Relative humidity: 20% to 85% (without dew condensation) Storage temperature: −25℃ to 65℃ Power supply: Rated voltage 100 V to 240 V AC, 50/60 Hz, 20 VA (max.) External dimensions: 96 (W) mm × 96 (H) mm × 115 (D) mm Case depth: approx. 105 mm (when panel-mounted) Mass: Approx. 550 g Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |