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	1
	PANEL-MOUNT TYPE CONDUCTIVITY METER (FOUR-WIRE TRANSMISSION, FOR WIDE RANGE)
	Features
· Simultaneous 4-channel output
Advanced concentration management is possible through simultaneous 3-channel output of low, middle, and high concentrations, etc.
· Five contact alarms
· Input signal can apply for analog output range selection
· Including nitric acid, phosphoric acid, sodium hydroxide, and 2 user-specified types, the concentration of a total of 5 types of chemical solutions can be output
· Built-in Communications Function
Technical specification
Measuring method: Four-electrode type
Cell constant: 0.1/cm
Temperature sensor specifications: Resistance thermometer bulb 1000 Ω/0℃
Measuring range: 
Conductivity: 0 mS/cm to 200 mS/cm (measuring range of conductivity before temperature compensation: 0 mS/cm to 500 mS/cm)
Temperature: 0.0℃ to 100.0℃
Display and output range:
Conductivity: 0 mS/cm to 2000 mS/cm
Temperature: −30℃ to 160℃
Concentration: 0% to 5%: NaOH, HNO3, H3PO4 0% to 100%: Concentration of the component
Temperature compensation range: 0℃ to 100℃ (a temperature less than 0℃ and more than 100℃ is calculated as an extension of the range)
Ambient temperature: −5℃ to 55℃
Relative humidity: 20% to 85% (without dew condensation)
Storage temperature: −25℃ to 65℃
Power supply: Rated voltage 100 V to 240 V AC, 50/60 Hz, 20 VA (max.)
External dimensions: 96 (W) mm × 96 (H) mm × 115 (D) mm
Case depth: approx. 105 mm (when panel-mounted)
Mass: Approx. 550 g
Documentation:
· CE marking and declaration conformity
· Documents for the Installation and Operational Qualification (IQ / OQ)
· Functional Specification - Design Qualification (DQ) and others relevant document
	pcs
	1



