| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | TOTAL ORGANIC CARBON MONITOR | Description suitable for the measurement ot total organic carbons in Purified Water (PW), Water For Injection (WFI) and Highly Purified Water (HPW). Features - UV-Oxidation and conductivity determination: This method eliminates the need for reagents and carrier gases. It delivers an accurate measurement at low concentrations, like for instance in WFI (Water For Injection) areas and reduces the cost of operation at the same time. - “Debubbler“ embedded: Bubbles in production systems often cause unstable readings in many instruments. The embedded “Debubbler“ eliminates the influence of air and guarantees a stable measurement. - Optional external ON/OFF control of contact input terminals. - Maintenance cycle: “ONLY ONCE PER YEAR“: Maintenance only consists of annual UV-lamp exchange and annual calibration. - Conforms to: United States Pharmacopeia: USP <643>. - Reference solutions for System Suitability Test (SST): 1,4 Benzoquinone and Sucrose - Choice of 3 installation methods: Possible to install on a pole, wall or desktop. Technical specification Measurement method: Continuous UV oxidation conductivity method Measurement limit: 0 ppb to 1000 ppb Measurement range: 0 ppb to 100/500/1000 ppb Repeatability: Within ±2% of full scale Display resolution: 0.1 ppb (0 ppb to 100 ppb/0 ppb to 500 ppb range) 1 ppb (0 ppb to 1000 ppb range) Transmission output: Number of output point: one 4 mA to 20 mA DC/0 V to 1 V DC (optional) insulated input/output Maximum load resistance: 600 Ω (minimum load resistance 100 kΩ (when optionally selected) Transmission output range: optional setting possible within measurement range Output target: TOC concentration Sample water condition Conductivity: less than 10 μS/cm (less than 2 μS/cm in 0 ppb to 100 ppb range) Temperature: 5℃ to 99℃ Pressure: 0.03 MPa to 0.50 MPa, Pressure fluctuation: within ±0.05 MPa (within the specified pressure range) Flow rate: 20 mL/min to 200 mL/min Environmental conditions Temperature: 5℃ to 40℃ Relative humidity: less than 85% Display function LCD (with a back light) TOC, conductivity, water temperature, alarm Power 100 V to 240 V AC ±10%, 50 Hz/60 Hz built-in fuse 250 V AC 3.15 A (T) Power consumption: 80 VA (max.) Outside dimension: 300 (W) × 400 (H) × 165 (D) mm Mass: 13 kg Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |