| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, FOR WIDE RANGE) | Features Tough - Responds to harsh conditions at the site process Operating temperature range(Transmitter)：-20 - 55℃ Rated power supply voltage: AC 100 - 240 V ± 10 ％ 50/60 Hz Robust aluminum die cast (Protection code：IP65) - CE Marking and FCC Compliance suitable for use at overseas sites ### Intelligence - Convenient function to display liquid temperature at the same time Transmission output available） - Sensor error can be detected immediately by self-diagnosis function - Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS ### Easy maintenance - Easy to view big display - Easy operation - Calibration log memory - Screw terminals that do not fall off are used in the terminal block for wiring Technical specification Measurement range Kaolin: 0 to 1000 (degrees) (Display range: 0 to 1100 (degrees)) or 0 to 1000 (mg/L) Formazin : 0 to 2000 (Display range: 0 to 2200 ) PSL: 0 to 100 (degrees) (Display range: 0 to 110 (degrees)) Transmission output setting range: Kaolin: Arbitrary setting within the range from 0.00 to 10.00 (degrees), 0.0 to 100.0 (degrees), 0 to 2200 (degrees) Display resolution 0.01 degrees (Degree range: 0 to 10 (degrees)) 0.1 degrees (Degree range: 0 to 100 (degrees)) 1 degrees (Degree range: 100 to 2200 (degrees)) Selectable between fixed range (decimal point) and automatic range changeover Performance Repeatability: Within ±2% of the reading or ±0.04 degrees, whichever is larger (checked using the solid span bottle) Linearity: The deviation at the mid-range point of the span calibration value is within ±2% of the span calibration value or ±0.04 degrees, whichever is larger Accuracy: Within ±2% of reading Transmission output Number of current outputs:2 Output type 4 mA to 20 mA DC, input/output isolation type Load resistance 900 Ω (max.) Linearity: Within ±0.08 mA (limited to the current linearity) Repeatability: Within ±0.02 mA (limited to the current repeatability) Error output: Provided with burnout function (3.8 mA or 21 mA) Hold function: Selectable between the latest value hold and the preset value hold Operating temperature range −20℃ to 55℃(no freezing) Operating humidity range: Relative humidity: 5% to 90% (no condensation) Storage temperature: −25℃ to 65℃ Power supply Rated power supply voltage: 100 V to 240 V AC ±10% 50/60 Hz Power consumption: 35 VA (max.) when an automatic cleaner is connected. Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |