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	FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, FOR WIDE RANGE)
	Features
Tough
· Responds to harsh conditions at the site process
Operating temperature range(Transmitter)：-20 - 55℃
Rated power supply voltage: AC 100  - 240 V ± 10 ％ 50/60 Hz
Robust aluminum die cast (Protection code：IP65)
· CE Marking and FCC Compliance suitable for use at overseas sites
Intelligence
· Convenient function to display liquid temperature at the same time Transmission output available）
· Sensor error can be detected immediately by self-diagnosis function
· Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS
Easy maintenance
· Easy to view big display
· Easy operation
· Calibration log memory
· Screw terminals that do not fall off are used in the terminal block for wiring
Technical specification
Measurement range
Kaolin: 0 to 1000 (degrees) (Display range: 0 to 1100 (degrees)) or 0 to 1000 (mg/L)
Formazin : 0 to 2000 (Display range: 0 to 2200 )
PSL: 0 to 100 (degrees) (Display range: 0 to 110 (degrees))
Transmission output setting range:
Kaolin: Arbitrary setting within the range from 0.00 to 10.00 (degrees), 0.0 to 100.0 (degrees), 0 to 2200 (degrees)
Display resolution	
0.01 degrees (Degree range: 0 to 10 (degrees))
0.1 degrees (Degree range: 0 to 100 (degrees))
1 degrees (Degree range: 100 to 2200 (degrees))
Selectable between fixed range (decimal point) and automatic range changeover
Performance
Repeatability: Within ±2% of the reading or ±0.04 degrees, whichever is larger (checked using the solid span bottle)
Linearity: The deviation at the mid-range point of the span calibration value is within ±2% of the span calibration value or ±0.04 degrees, whichever is larger
Accuracy: Within ±2% of reading
Transmission output	
Number of current outputs:2
Output type	4 mA to 20 mA DC, input/output isolation type
Load resistance	900 Ω (max.)
Linearity: Within ±0.08 mA (limited to the current linearity)
Repeatability: Within ±0.02 mA (limited to the current repeatability)
Error output: Provided with burnout function (3.8 mA or 21 mA)
Hold function: Selectable between the latest value hold and the preset value hold
Operating temperature range	−20℃ to 55℃(no freezing)
Operating humidity range: Relative humidity: 5% to 90% (no condensation)
Storage temperature: −25℃ to 65℃
Power supply
Rated power supply voltage: 100 V to 240 V AC ±10% 50/60 Hz
Power consumption: 35 VA (max.) when an automatic cleaner is connected.
Documentation:
· CE marking and declaration conformity
· Documents for the Installation and Operational Qualification (IQ / OQ)
· Functional Specification - Design Qualification (DQ) and others relevant document
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