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	FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, IMMERSION TYPE)
	Features
Tough
· Responds to harsh conditions at the site process
Operating temperature range(Transmitter)：-20 - 55℃
Rated power supply voltage: AC 100  - 240 V 50/60 Hz
Robust aluminum die cast (Protection code：IP65)
· CE Marking and FCC Compliance suitable for use at overseas sites
· Long-life infrared LED light source
 
Intelligence
· Sensor error can be detected immediately by self-diagnosis function
· Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS
· Turbidity or SS can be selected for measurement
· Function that automatically corrects changes in the light intensity of the light source enables stable measurement
 
Easy maintenance
· Easy to view big display
· Easy operation
· Calibration log memory
· Screw terminals that do not fall off are used in the terminal block for wiring
· Built-in proprietary electric rotary wiper cleaner resistant to contamination
Technical specification
· Measurement method
Right angle scattering/ transmission: Measuring principle suitable for full range turbidity and SS within 200 mg/L by Kaorin
Transmission: Measuring principle suitable for SS over 50 mg/L by Kaorin
· Measurement range
Turbidity: Scattering/transmission: Formazin 0 NTU to 4000 NTU Kaolin: 0 to 2000 (Display range: 2001 to 4000)
SS: Scattering/ transmission Kaolin 0 mg/L to 2000 mg/L (Display range: 2001 mg/L to 4000 mg/L)
· Repeatability
Turbidity Scattering/ transmission
Formazin: 0 NTU to 2000 NTU Within ±2% of the reading or ±0.5
NTU, whichever is larger 2001 NTU to 4000 NTU	Within ±3% of the reading
Kaolin 0 to 1000 Within ±2% of the reading or ±0.5, whichever is larger
1001 to 2000 Within ±5% of the reading
SS Scattering/ transmission
Kaolin 0 mg/L to 1000 mg/L Within ±2% of the reading or ±0.5mg/ L, whichever is larger 1001 mg/L to 2000 mg/L Within ±5% of the reading
Transmission :0 mg/L to 2000 mg/L Within ±5% of the reading or ±0.5mg/ L, whichever is larger
· Transmission output
Number of current outputs: 2
Output type: 4 mA to 20 mA DC, input/output isolation type
Load resistance	900 Ω (max.)
Linearity: Within ±0.08 mA (limited to the current linearity)
Repeatability: Within ±0.02 mA (limited to the current repeatability)
Error output: Provided with burnout function (3.8 mA or 21 mA)
Transmission holding: One of the latest value or preset value can be specified
· Operating temperature range: −20℃ to 55℃
· Operating humidity range: Relative humidity: 5% to 90% (no condensation)
· Storage temperature: −25℃ to 65℃
· External dimensions: 180 (W)  × 155 (H)  × 115 (D) mm (excluding brackets)
· Mass :Mainframe: approx. 3.5 kg
· Hood, bracket: approx. 1 kg
Documentation:
· CE marking and declaration conformity
· Documents for the Installation and Operational Qualification (IQ / OQ)
· Functional Specification - Design Qualification (DQ) and others relevant document
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