| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, IMMERSION TYPE) | Features Tough - Responds to harsh conditions at the site process Operating temperature range(Transmitter)：-20 - 55℃ Rated power supply voltage: AC 100 - 240 V 50/60 Hz Robust aluminum die cast (Protection code：IP65) - CE Marking and FCC Compliance suitable for use at overseas sites - Long-life infrared LED light source Intelligence - Sensor error can be detected immediately by self-diagnosis function - Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS - Turbidity or SS can be selected for measurement - Function that automatically corrects changes in the light intensity of the light source enables stable measurement Easy maintenance - Easy to view big display - Easy operation - Calibration log memory - Screw terminals that do not fall off are used in the terminal block for wiring - Built-in proprietary electric rotary wiper cleaner resistant to contamination Technical specification - Measurement method Right angle scattering/ transmission: Measuring principle suitable for full range turbidity and SS within 200 mg/L by Kaorin Transmission: Measuring principle suitable for SS over 50 mg/L by Kaorin - Measurement range Turbidity: Scattering/transmission: Formazin 0 NTU to 4000 NTU Kaolin: 0 to 2000 (Display range: 2001 to 4000) SS: Scattering/ transmission Kaolin 0 mg/L to 2000 mg/L (Display range: 2001 mg/L to 4000 mg/L) - Repeatability Turbidity Scattering/ transmission Formazin: 0 NTU to 2000 NTU Within ±2% of the reading or ±0.5 NTU, whichever is larger 2001 NTU to 4000 NTU Within ±3% of the reading Kaolin 0 to 1000 Within ±2% of the reading or ±0.5, whichever is larger 1001 to 2000 Within ±5% of the reading SS Scattering/ transmission Kaolin 0 mg/L to 1000 mg/L Within ±2% of the reading or ±0.5mg/ L, whichever is larger 1001 mg/L to 2000 mg/L Within ±5% of the reading Transmission :0 mg/L to 2000 mg/L Within ±5% of the reading or ±0.5mg/ L, whichever is larger - Transmission output Number of current outputs: 2 Output type: 4 mA to 20 mA DC, input/output isolation type Load resistance 900 Ω (max.) Linearity: Within ±0.08 mA (limited to the current linearity) Repeatability: Within ±0.02 mA (limited to the current repeatability) Error output: Provided with burnout function (3.8 mA or 21 mA) Transmission holding: One of the latest value or preset value can be specified - Operating temperature range: −20℃ to 55℃ - Operating humidity range: Relative humidity: 5% to 90% (no condensation) - Storage temperature: −25℃ to 65℃ - External dimensions: 180 (W) × 155 (H) × 115 (D) mm (excluding brackets) - Mass :Mainframe: approx. 3.5 kg - Hood, bracket: approx. 1 kg Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |