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	FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, HIGH-SENSITIVITY)
	Features
Tough
· Responds to harsh conditions at the site process
　　Operating temperature range(Transmitter)：-20 - 55℃
　　Rated power supply voltage: AC 100  - 240 V ±10％ 50/60 Hz
　　Robust aluminum die cast (Protection code：IP65)
· CE Marking and FCC Compliance suitable for use at overseas sites
· Long-life infrared LED light source
 
Intelligence
· Sensor error can be detected immediately by self-diagnosis function
· Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS
· Blinking light sources that can eliminate ambient light
 
Easy maintenance
· Easy to view big display
· Easy operation
· Calibration log memory
· Screw terminals that do not fall off are used in the terminal block for wiring
Technical specification
· Measurement range
Kaolin	0 to 10.00 (degrees) (Display range: 0 to 11.00 (degrees))
Formazin (NTU)	0 to 10.00 (NTU) (Display range: 0 to 11.00 (NTU))
PSL 0 to 10.00 (degrees) (Display range: 0 to 11.00 (degrees))
· Performance
Repeatability: Within ±2% of the reading or ±0.04 degrees, whichever is larger (checked using the span bottle)
Linearity: The deviation at the mid-range point of the span calibration value is within ±2% of the span calibration value or ±0.04 degrees, whichever is larger
· Transmission output	
Number of current outputs: 2
Output type: 4 mA to 20 mA DC, input/output isolation type
Load resistance	900 Ω(max.)
Linearity: Within ±0.08 mA (limited to the current linearity)
Repeatability: Within ±0.02 mA (limited to the current repeatability)
Error output: Provided with burnout function (3.8 mA or 21 mA)
Hold function: Selectable between the latest value hold and the preset value hold
· Operating temperature range: −20℃ to 55℃ (no freezing)
· Operating humidity range: Relative humidity: 5 to 90 (%) (no condensation)
· Storage temperature: −25℃ to 65℃
· Power supply: Rated power supply voltage 100 V to 240 V AC ±10% 50/60 Hz  Power consumption: 35 VA (max.) when an automatic cleaner is    connected.
· External dimensions: 180 (W)  × 155 (H)  × 115 (D) mm (excluding brackets)
· Mass: Mainframe: approx. 3.5 kg (excluding brackets), hood, bracket: approx. 1 kg
Documentation:
· CE marking and declaration conformity
· Documents for the Installation and Operational Qualification (IQ / OQ)
· Functional Specification - Design Qualification (DQ) and others relevant document
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