| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | FIELD-INSTALLATION TYPE TURBIDITY METER (FOUR-WIRE TRANSMISSION, HIGH-SENSITIVITY) | Features Tough - Responds to harsh conditions at the site process Operating temperature range(Transmitter)：-20 - 55℃ Rated power supply voltage: AC 100 - 240 V ±10％ 50/60 Hz Robust aluminum die cast (Protection code：IP65) - CE Marking and FCC Compliance suitable for use at overseas sites - Long-life infrared LED light source Intelligence - Sensor error can be detected immediately by self-diagnosis function - Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS - Blinking light sources that can eliminate ambient light Easy maintenance - Easy to view big display - Easy operation - Calibration log memory - Screw terminals that do not fall off are used in the terminal block for wiring Technical specification - Measurement range Kaolin 0 to 10.00 (degrees) (Display range: 0 to 11.00 (degrees)) Formazin (NTU) 0 to 10.00 (NTU) (Display range: 0 to 11.00 (NTU)) PSL 0 to 10.00 (degrees) (Display range: 0 to 11.00 (degrees)) - Performance Repeatability: Within ±2% of the reading or ±0.04 degrees, whichever is larger (checked using the span bottle) Linearity: The deviation at the mid-range point of the span calibration value is within ±2% of the span calibration value or ±0.04 degrees, whichever is larger - Transmission output Number of current outputs: 2 Output type: 4 mA to 20 mA DC, input/output isolation type Load resistance 900 Ω(max.) Linearity: Within ±0.08 mA (limited to the current linearity) Repeatability: Within ±0.02 mA (limited to the current repeatability) Error output: Provided with burnout function (3.8 mA or 21 mA) Hold function: Selectable between the latest value hold and the preset value hold - Operating temperature range: −20℃ to 55℃ (no freezing) - Operating humidity range: Relative humidity: 5 to 90 (%) (no condensation) - Storage temperature: −25℃ to 65℃ - Power supply: Rated power supply voltage 100 V to 240 V AC ±10% 50/60 Hz Power consumption: 35 VA (max.) when an automatic cleaner is connected. - External dimensions: 180 (W) × 155 (H) × 115 (D) mm (excluding brackets) - Mass: Mainframe: approx. 3.5 kg (excluding brackets), hood, bracket: approx. 1 kg Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |