| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | WATER DISTRIBUTION MONITOR | Features - Equipped with sensors to monitor four water quality parameters. Optional sensors can be added to monitor as many as seven water quality parameters. No reagents are required for the measurement. - The instrument is operated with a touch screen panel for easy operation and display of measured values. - The compact size allows for more installation location options and space for operation. Technical specification Parameters: Turbidity, color, free chlorine, water pressure (standard); pH, conductivity Range: Turbidity: 0 to 2 FTU, 0 to 4 FTU Color: 0 to 10 CU, 0 to 20 CU Free chlorine: 0 to 2 mg/L Water pressure: 0 to 1 MPa Repeatability: Turbidity: Full scale ±2.5% Color: Full scale ±5.0% Free chlorine: Full scale ±2.5% Water pressure: Full scale ±1.0% Display: LCD display: 320 x 240 dot B&W graphic LCD with backlight (touch panel type) Standard version simultaneously displays four measurement categories (turbidity, color, free chlorine, and water pressure); Optional version simultaneously displays up to seven parameters (the four parameters plus pH, conductivity, and water temperature) Automatic zero calibration (Turbidity, color, free chlorine) Calibration method: Measurement of filtered sample water Trigger: Internal control (timer) or external control. Interval: 1 to 9999 hours (can be set as desired within that range) Time: 20~99 minutes(can be set as desired within that range) Automatic cleaning (Turbidity, color) Cleaning method: Wiper cleaning Trigger: Internal control (timer) or external control. Interval: 5 to 9999 hours (can be set as desired within that range) (Free chlorine) Continuous cleaning with beads grinder Sample water conditions Temperature: 0 to 40°C (must not freeze) Pressure: 0.1 to 0.75 Mpa Flow rate: over 100 mL/min Flow rate for measurement: 50 to 100 mL/min Ambient temperature/humidity: Temperature; 0 to 40°C; Humidity: 85% or less Power: AC 100 V to AC 230 V, ±10%, 50/60 Hz Power consumption: AC 100 V to AC 120 V: 120 VA; AC 100 V to AC 230 V: 150 VA Mass Approx. 15 kg (Measurement unit only) External dimensions: 350 mm (W) x 160 mm (D) x 420 mm (H) Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |