| Item No | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | Field-installation Type pH Meter (Two-Wire Transmission) | Description measures pH and temperature in the sample water. Suitable for monitoring and control of various production processes and wastewater treatment, etc. Many types of electrodes are available to match the sample properties. Features Measurement range pH: 0 - 14 Temperature: 0 - 100 ℃ Tough - Responds to harsh conditions at the site process Operating temperature range(Transmitter)：-20 - 60℃ Rated voltage: DC 24 Robust aluminum die cast (Protection code：IP65) - CE Marking and FCC Compliance suitable for use at overseas sites Intelligence - Convenient function to display liquid temperature at the same time（Transmission output available） - Sensor error can be detected immediately by self-diagnosis function - Designed to be environmentally friendly (Pb-free) in response to WEEE and RoHS Easy maintenance - Easy to view big display - Easy operation - Calibration log memory - Screw terminals that do not fall off are used in the terminal block for wiring Technical specification Combined sensor: Glass electrode Measurable range: pH : pH0 to pH14 (display range: pH−1 to pH15) Temperature: 0℃ to 100℃ Display range when automatic detection capability for each temperature sensor type is used: −10℃ to 110℃ Display range when temperature sensor type is manually set: −20℃ to 130℃ Display resolution pH: 0.01 pH Temperature: 0.1℃ Performance pH: Repeatability Within ±0.03 pH (response for equivalent input) Linearity Within ±0.03 pH (response for equivalent input) Temperature Repeatability ±0.3℃ (response for equivalent input) Linearity ±0.3℃ (response for equivalent input) Transmission output Output type: 4 mA to 20 mA DC: input/output isolated type Load resistance Maximum: 600Ω Case of 24 V DC power supply (*Note 1) Linearity: Within ±0.08 mA (output only) Repeatability: Within ±0.08 mA (output only) Output range: pH: Selection from preset ranges or free range input within measuring range. Occasional out for error Hold or burnout to either 3.8 mA or 21 mA Transmission hold: In the maintenance mode, transmission signal is held at the latest value or preset value. In the calibration mode, transmission signal can be alive or held. Operating temperature range: −20℃ to 60℃ (without freeze) Operating humidity range: Relative humidity: 5% to 90% (without condensation) Storage temperature−25℃ to 65℃ Power supply: Rated voltage 24 V DC (operating voltage range: 21 V to 32 V DC) (*Note 1) Power consumption: 0.6 W max. Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | Pcs | 1 |