| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | Panel-mount Type pH Meter (Four-Wire Transmission) | Description measures pH and temperature in the sample water. Suitable for monitoring and control of various production processes and wastewater treatment, etc. Many types of electrodes are available to match the sample properties. Features - Easy pH correction After immersing the pH poles in the standard fluid, simply follow the instructions displayed and press the ENT key. The automatic correction process is plain and simple. - Automatic diagnosis function for faults in electrode properties If there is an asymmetry potential, sensitivity or response speed in correction, the HP-480 automatically determines whether there is a fault in the pH electrodes, and displays an error message. - Compatible with five types of temperature compensation element Five types of temperature compensation element are supported: 500 Ω, 6.8 kΩ, 1 kΩ, 10 kΩ and 350 Ω. - A line-up of lead-free pH electrodes Uses lead-free glass. In addition, almost all sections of the pH electrodes, including lead wires and soldered junctions are lead-free. - Uses lead-free solder To reduce the burden on the environment, lead-free solder has been used to mount the components on the printed circuit board. Technical specification Measuring range: 0 pH to 14 pH: 0.01 pH resolution. Temperature 0℃ to 100℃: 1℃ resolution (selectable display). Transmission output: 4 mA to 20 mA DC: input/output isolated type Maximum load resistance: 900 Ω Transmission output range: Free Repeatability and linearity: ±0.05 pH, ±0.08 mA (for transmission output range of 0 pH to 14 pH) Contact output Outputs: 2 points Alarm contact output (R1 and R2) Contact type: relay contact, SPDT (1c) Contact rating: 240 V AC, 3 A and 30 V DC, 3 A (resistance load) Contact function: selectable from upper/lower limit operation (ON/OFF control), alarm, and maintenance. Calibration function: Two-point automatic calibration or manual calibration. Two-point automatic calibration: with automatic validation of the electrode characteristics. Types of standard solution: pH 2, pH 4, pH 7, pH 9 and pH 10 (JIS) Combination of standard solutions: pH 7 and one of the others Manual calibration: Freely selectable, but the difference should be over 2 pH. Temperature calibration (one point). Temperature compensation element: Selectable from compensation ON (350Ω (25℃), 500Ω (25℃),1 kΩ (0℃)6.8 kΩ (25℃) and 10 kΩ (25℃)) and compensation OFF. Temperature- compensated range: 0℃ to 100℃ Ambient temperature: −5℃ to 45℃ Relative humidity: 20% to 85% (without dew condensation) Storage temperature: −25℃ to 65℃ Power supply: Rated voltage 100 V AC to 240 V AC, 50 Hz/60 Hz, 10 VA (max.) External dimensions: 48 (W) × 96 (H) × 115 (D) mm Case depth: approx. 105 mm Mass: Approx. 400 g Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |