| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | Panel-mount Type pH Meter (Four-Wire Transmission, 4 point alarm, Time sharing Proportional Control) | Description High-end pH meter with a four-point alarm and a time sharing proportional control function. It connects pH sensor and measures pH and temperature in the sample water. It can also be linked with a cleaning device via an external input, extending its automatic operation time. Many types of electrodes are available to match the sample properties. Features - Time sharing proportional control function with F Zone Unlike commonly used time sharing proportional control functions, whose on-off frequency has generally been fixed, the HP-960FTP includes an control method called F Zone, whereby as the pH difference becomes smaller, it changes not just the time for which the pump injecting pH regulator is on, but also its frequency. - Four control outputs equipped You can set values of your choice for the upper and lower limits of the four control outputs, and switch them to either on-off control or time sharing proportional control. - Automatic diagnosis function for faults in electrode properties If there is an asymmetry potential, sensitivity or response speed in correction, the HP-960FTP automatically determines whether there is a fault in the pH electrodes, and displays an error message. - Can be linked with a cleaning device via an external input An input signal from a cleaning device can be used to put the output of the measurement value on hold. - Compatible with five types of temperature compensation elements Five types of temperature compensation element are supported: 500 Ω, 6.8 kΩ, 1 kΩ, 10 kΩ and 350 Ω. - Easy pH correction After immersing the pH electrodes in the standard fluid, simply follow the instructions displayed and press the ENT key. The automatic correction process is plain and simple. - A line-up of lead-free pH electrodes Uses lead-free glass. In addition, almost all sections of the pH electrodes, including lead wires and soldered junctions are lead-free. - Uses lead-free solder To reduce the burden on the environment, lead-free solder has been used to mount the components on the printed circuit board. Technical specification Measuring range: 0 pH to 14 pH: 0.01 pH resolution. Temperature 0℃ to 100℃: 1℃ resolution (selectable display). Repeatability: Within±0.05 pH (Equivalent input) Linearity: Within±0.05 pH (Equivalent input) Transmission output: 4 mA to 20 mA DC, input/output isolated type Maximum load resistance: 900 Ω Transmission output range: Free Contact output: Outputs: 4 points Alarm contact output (R1, R2, R3, and R4) Contact type: relay contact, SPDT (1c) Contact rating: 240 V AC 3 A and 30 V DC 3 A (resistance load) Contact function: selectable from upper/lower limit operation (ON/OFF control), alarm, and maintenance. Temperature compensated range: 0℃ to 100℃ Ambient temperature: −5℃ to 45℃ Relative humidity: 20% to 85% (without dew condensation) Storage temperature: −25℃ to 65℃ Power supply: Rated voltage 100 V to 240 V AC, 50 Hz/60 Hz, 10 VA (max.) External dimensions: 96 (W) × 96 (H) × 115 (D) mm Case depth: approx. 105 mm Mass: Approx. 550 g Documentation: - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |