| Item No. | Item Name | Technical Specifications | Unit of Measurement | Quantity |
| --- | --- | --- | --- | --- |
| 1 | ELECTRON MICROSCOPE ANALYZERS | Features - Extremely fast mapping using only moderate beam currents - Analysis of beam-sensitive materials at extremely low beam currents (< 10 pA), e.g. biological or semiconductor samples - Investigation of samples with topography, avoiding shadowing effects - Analysis of nanoparticles and nanostructures at low kV and highest magnification - Measurement of thin samples (e.g. TEM lamellae) and other specimens with low X-ray yield in the SEM. Application Advanced Elemental Distribution Mapping of Life Science Samples Automated Deconvolution and Quantification of Peak Overlaps in MoS2 Elemental Mapping of Niobium in Engineering Steel Technical specification Energy resolution: Energy resolution of 126 eV at Mn Ka, 51 eV at C Ka, 60 eVat F Ka Also available: 129 eV at Mn Ka, 58 eV at C Ka, 66 eV at F Ka 133 eV at Mn Ka, 65 eV at C Ka, 73 eV at F Ka Stated in compliance with ISO 15632:2012 and guaranteed at 100,000 cps input count rate Element range: Boron (5) to americium (95), further details see table below Input count rate: Up to 4,000,000 counts per second (cps) Output count rate: Up to 1,600,000 cps Active area :60 mm2 (4 × 15 mm2) Cooling: Liquid nitrogen-free Peltier cooler, vibration-free Operating pressure: Max. 30 Pa permissible during detector operation Detector mount: Horizontal port with 33 mm minimum diameter required (fixed Y and Z position), larger diameter to make use of detector XYZ adjustment Included Articles Window changer: 3-position window changer with 2 polymer windows and vacuum passthrough, see also table below Interference: Minimal interference with SEM through use of non-magnetic materials Documentation - CE marking and declaration conformity - Documents for the Installation and Operational Qualification (IQ / OQ) - Functional Specification - Design Qualification (DQ) and others relevant document | pcs | 1 |